## Slide 1: Overview of Next Generation Short Read Sequencing Technologies
Elaine R. Mardis, Ph.D., FAACR
Co-Executive Director, The Steve and Cindy Rasmussen Institute for Genomic Medicine at Nationwide Children’s Hospital
The Steve and Cindy Rasmussen Nationwide Foundation Endowed Chair in Genomic Medicine
Professor of Pediatrics and Neurosurgery, The Ohio State University College of Medicine

## Slide 2: Next Generation Sequencing Basics
How massively parallel sequencing works

## Slide 3: NGS has transformed biomedical inquiry
E.R. Mardis, Nature (2011) 470: 198-203, Ann. Rev. Analyt. Chem. (2013)

## Slide 4: Next Generation Sequencing: the basics
  NGS library construction combines DNA or RNA fragments with custom synthetic DNA adapters by ligation or transposon insertion
  The resulting library fragments are amplified on a solid support (either a bead, flat surface or nanowell-covered surface) with covalently attached adapters complimentary to the library adapters
  Sequencing reactions couple nucleotide incorporation and detection in a step-wise fashion, detecting hundreds of millions to billions of sequencing reactions per instrument run = “massively parallel sequencing”
  Shorter read lengths than capillary sequencers = requires specialized bioinformatics-based analyses

## Slide 5: Input DNA for NGS Libraries
Generally speaking, NGS sequencing libraries are ‘short insert’, with fragment sizes ranging from ~200-600 bp
Starting material for NGS libraries may be derived from numerous different sources, including:
- high molecular weight genomic DNA (cell lines, blood, fresh or frozen tissue, buccal swab/cheek scrape)
- PCR products, including from multiplex PCR
- low molecular weight/degraded DNA (formalin-fixed paraffin-embedded (FFPE) tissues, forensic specimens, Neanderthal bones)
So, the first consideration is the method of DNA isolation needed, and then evaluating the quality/intactness and yield of DNA that results

## Slide 6: Evaluating Input DNA Quality and Quantity
DNA Quality/Intactness:
- cheapest = agarose gel and ethidium staining
- easiest = precast gel (Flash)
- large sample numbers: Agilent BioAnalyzer or TapeStation
DNA Quantity:
- NanoDrop (with caveats)
- Qubit
- PicoGreen + 96 well plate reader
If high mw and concentration, fragmentation is next
If low mw and concentration, challenges and adjustments will be required!

## Slide 7: Fragmenting High MW Genomic DNA
Mechanical methods:
- syringe/needle high pressure shearing device
- HydroShear device
- Covaris ultrasonicator or G-tube device
Enzymatic methods:
- restriction enzyme(s) cleavage
- Zinc treatment
- proprietary nuclease cocktail / “fragmentase”
Initially, evaluate one or more post-shearing samples on gel for size range optimization

## Slide 8: Transposon-based Library Construction

## Slide 9: Adapter Ligation-based Library Construction
Shear or enzymatically treat high molecular weight DNA to fragment, OR derive DNA fragments from other preparatory methods
Enzymatic treatments to blunt ends
Ligate synthetic DNA adapters (with indexes), PCR amplify
Quantitate library
Proceed to WGS, or isolate exome or specific panel by hybrid capture

## Slide 10: PCR-related Problems in NGS
PCR is an effective vehicle for amplifying DNA, however…
In NGS library construction, PCR can introduce preferential amplification (“jackpotting”) of certain fragments
    Duplicate reads with exact start/stop alignments
    Need to “de-duplicate” after alignment and keep only one pair
    Low input DNA amounts favor jackpotting due to lack of complexity in the fragment population
PCR also introduces false positive artifacts due to substitution errors by the polymerase
    If substitution occurs in early PCR cycles, error appears as a true variant
    If substitution occurs in later cycles, error typically is drowned out by correctly copied fragments in the cluster
Cluster formation is a type of PCR (“bridge amplification”)
    Introduces bias in amplifying high and low G+C fragments
    Reduced coverage at these loci is a result

## Slide 11: NGS Library Multiplexing and Molecular Barcoding
As throughput on NGS sequencers has increased, the ability to pool libraries or reaction products together is needed
DNA barcoding schemes permit the addition of specific sequences to each library, enabling equimolar pooling (“multiplexing”)
Post-run de-multiplexing is accomplished by software that bins reads sharing the same barcodes or barcode combinations
As sequencing costs permit deeper sequencing of libraries, the use of unique molecular identifiers (UMIs) has become necessary to separate true mutations from sequencing “noise” during data analysis

## Slide 12: Unique Molecular Identifiers: UMI Barcoding
UMI in i7 arm of adapter is
added during ligation
Ligation product with i7 UMI

## Slide 13: Dual Functionality of UMIs
De-duplication
Polymerase or base-calling error elimination

## Slide 14: Illumina Sequencing Basics
Sequencing by Synthesis

## Slide 15: Cluster Amplification on FlowCell
Quantitating library fragments is an essential first step to cluster formation
Diluted library fragments are introduced to flow cell, and are amplified in situ using the covalently attached complementary adapter sequences
Cluster amplification is required to produce sufficient signal for detecting the sequencing reaction results at each nucleotide addition step = each cluster is derived from a single library fragment

## Slide 16: Sequencing Chemistry for SBS
Excitation
Emission
Incorporate
Detect
De-block
Cleave fluor

## Slide 17: Sequencing by Synthesis

## Slide 18: Patterned Flow Cell

## Slide 19: Four-, Two- and One-Color Chemistry
The MiSeq and HiSeq Series Systems currently use four-channel SBS.
The MiniSeq, NextSeq, and NovaSeq Systems use two-channel chemistry.
The iSeq 100 use one-channel chemistry.

## Slide 20: Illumina Platforms
High accuracy reads, paired end reads, range of capacity and throughput
Longer read lengths on some platforms (MiSeq)
Largest user base with accompanying software applications, Dragen rapid alignment/analysis

## Slide 21: Ion Torrent Sequencing Technology
Sequencing by pH Sensing

## Slide 22: pH Sensing of Base Incorporation
On-bead fragment amplification by emulsion PCR
Unlabeled native dNTPs
Undefined read lengths

## Slide 23: Platforms
Low substitution error rate, in/dels problematic, single end reads
Inexpensive and fast turn-around for data production
Clinical platform (Genexus)
Microbiome, immune gene sequencing applications
IonTorrent Genexus
Ion Torrent S5
Ion Chef

## Slide 24: Sub-setting the Genome for NGS-based Assays
Exome and Gene Panels

## Slide 25: Direct Genomic Selection

## Slide 26: Solution-Phase Hybrid Capture
Hybrid capture - specific sequences from a whole genome library are selected by hybridization to probes that correspond to human exons or gene panels.
Probe DNAs are biotinylated, so hybrids can be removed from solution with streptavidin magnetic beads.
An “exome” by definition, is captured with probes corresponding to the exons of all genes annotated in the reference genome.
Custom capture reagents or “gene panels” can be synthesized to target specific loci that may be of clinical interest.

## Slide 27: Hybrid Capture Terminology
Notes: BED File!!

## Slide 28: Coverage Definitions for Hybrid Capture

## Slide 29: Evaluating Coverage Post-Alignment and De-dup

## Slide 30: Evaluating Probe Coverage Breadth at Fixed Depth

## Slide 31: Probe Coverage and Outlier Analysis

## Slide 32: Hybrid Probe Placement Matters: Missed Splice Site Mutation
O.L. Griffith et al., Nat. Commun. 2018

## Slide 33: Data visualization
Examining NGS data

## Slide 34: Integrated Genomics Viewer (IGV)
http://www.broadinstitute.org/igv

## Slide 35: IGV Visualization of Tumor vs. Normal : Somatic SNV

## Slide 36: IGV: Somatic Insertion/Deletion

## Slide 37: Copy Number Alterations Genome-wide

## Slide 38: Copy Number Alterations: Arm-level and Focal

## Slide 39: RNAseq

## Slide 40: RNAseq: Why
Increasing emphasis on sequencing the transcriptome, combined with multiple types of computational analysis
Fusion detection, exon skipping, allele-specific silencing, epigenetic links to driver gene amplification, immune deconvolution (in cancer)
Multiple sizes of RNAs are being identified and linked causally to disease and other mechanisms of gene regulation (miRNAs, lncRNAs)
Analysis identifies differential gene expression, alternative splicing, RNA editing
Provides a gene set for organisms without a reference genome

## Slide 41: RNAseq “Gotchas”
RNA is sensitive to degradation from heat and FFPE treatment
Genomic DNA can co-isolate with RNA and cause multiple problems
Isolating mRNA decreases the total amount of RNA into library prep
Strandedness is important
Short reads make detecting certain RNA-specific attributes or aberrations difficult (exon skipping, etc.)

## Slide 42: QC of RNA is Essential

## Slide 43: RNAseq Overview

## Slide 44: rRNA Depletion using RiboZero Reagent

## Slide 45: NEBNext Ultra II Directional
Total RNA
(100 ng – 1 µg)
rRNA depletion
RNA fragmentation
First-strand cDNA synthesis
Second-strand cDNA synthesis
End Prep
Adapter ligation
(IDT xGen Dual Index UMI Adapters)
Enrichment PCR
Library QC
Qubit DNA Concentration
Bioanalyzer/TapeStation visualization
KAPA Library Quant qPCR
Anneal to DNA probes
Digest DNA/RNA hybrid with RNase H
Digest DNA probes with DNase
SPRIselect purification
SPRIselect purification
SPRIselect purification
SPRIselect purification
Repair DNA ends
3’ adenylation
qPCR cycle
optimization
Hybrid Capture
IDT xGen Lockdown Probes

## Slide 46: RNA “stranded” Sequencing

## Slide 47: Conclusions
Widespread use of NGS platforms for biological research has changed the scale of biomedical inquiry and discovery
A variety of platforms and approaches are available for implementation of NGS, each with strengths and weaknesses.
Due to the size and complexity of the resulting data sets, validated approaches to assure quality of the data, along with analytical pipelines having appropriate sensitivity, specificity and reproducibility are required.
Similarly, data visualization interfaces are needed to permit genome-wide and locus-specific evaluation of variants and underlying data support.